Clio gun

I measure

Rp			     Ti			Rk		Triode

H.V.

Rb

Cl					Rg	           Cg

U measure

## HV = 90kV	Cl = 2 nF	Rp = 1k	Rk = 50	Rg = 45 k/2A	Cg = 30pF

Ti => Curent transformer 1V/amp 200MHz

# Rb => HV voltage divider 0.01% DC 0.5% AC up to 10MHz

U/U = 7,8.10-4 for 2 Amps

W = 8,1 J

IHV = 3,8 mA

PHV = 340 W

CTF3 gun

I measure

Rp			     Ti			Rk		Triode

H.V

Rb

Rg	           Cg

U measure

## HV=90/150 kV		Rp=     Rk=	Rg=12/20 k/7A		Cg=120/80 nF

Ti => Curent transformer 1V/amp 200MHz

# Rb => HV voltage divider 0.01% DC 0.5% AC up to 10MHz

U/U = 1.10-3 for 7 Amps

HV = 90kV				HV = 150kV

W = 486 J					W = 900 J

IHV = 23 mA				IHV = 23 mA

PHV = 2 k W				PHV = 3,4 k W

Interface remote control

Manual Control

# GUN

Control

Interface

# VME

Interface

Electronics

. Control interface has to manage:

.Specificities of the gun

. Remote control

. Communication with VME

. Hardware:

. G64, CPU, in/out analog and logical cards

. Integrated Automation  (Simatic)

Tests of Guns

Tests at LAL:

# GUN

Manual Control

Electronics

. Electronics tests.

. H.V. Tests

. Conditionning

. Beam tests (shape, stability, accuracy)

. Behavior of electronics during voltage break downs

Tests at CERN:

Control

Interface

# VME

Interface

Gun simulator

. Tests of interface and control

Sensitivity of beam measurements

Beam monitors :

. ECM (Electrode Capacitive de Mesure) => 50V/mA

Z=50 for Clio gun

Z=1M for CTF3 gun (has to be tested)

. WCM (Wall Current Monitor)  => 2,8V/A

Z=50

Low frequency cut off (f-3dB) = 100KHz

High frequency cut off (f-3dB) = 1,5GHz  ( 280ps FWHM)